## Blad1
MEASURING FUNCTION |  | MEASURING ELEMENT | [5 empty] | rdl:RDS222444 |  | A <MEASURING ELEMENT> is a <MEASURING FUNCTION> used for the measurement of a <Property>. It is defined independently from any other <INSTRUMENTATION DEVICE> that is handling that measurement result and that may or may not house that <MEASURING ELEMENT> | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | ACCELERATION MEASURING ELEMENT | [5 empty] | rdl:RDS222500 |  | An <ACCELERATION MEASURING ELEMENT> is a <MEASURING ELEMENT> for <ACCELERATION> measurement | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | ACCELERATION MEASURING ELEMENT |  |  | ACCELEROMETER | [5 empty] | rdl:RDS2225310 |  | An <ACCELEROMETER> is an <INSTRUMENTATION ITEM> and an <ACCELERATION MEASURING ELEMENT> for the measurement of <ACCELERATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | ANALYSIS MEASURING ELEMENT | [5 empty] | rdl:RDS222501 |  | An <ANALYSIS MEASURING ELEMENT> is a <MEASURING ELEMENT> for <ANALYSIS> purposes | [5 empty] | dm:ClassOfFunctionalObject
ANALYSIS MEASURING ELEMENT |  |  |  | CONDUCTIVITY MEASURING ELEMENT | [5 empty] | rdl:RDS2225311 |  | A <CONDUCTIVITY MEASURING ELEMENT> is an <> for the measurement of <CONDUCTIVITY(electrolytic)>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | CONDUCTIVITY MEASURING ELEMENT |  |  | CONDUCTIVITY CELL | [5 empty] | rdl:RDS5770462 |  | An <CONDUCTIVITY CELL> is an <INSTRUMENTATION ITEM> and a <CONDUCTIVITY MEASURING ELEMENT> that is, when it has a Cell Constant K = 1.0, formed by two 1 cm^2 surfaces spaced 1 cm apart. Cells of different physical configuration are characterized by their own cell constant. This cell constant is a function of the electrode areas, the distance between the electrodes and the electrical field pattern between the electrodes. | [5 empty] | dm:ClassOfInanimatePhysicalObject
CONDUCTIVITY CELL |  |  |  |  | CONDUCTIVITY CELL - CONTACTING | [5 empty] | rdl:RDS222502 |  | A <CONDUCTIVITY CELL - CONTACTING> is a <CONDUCTIVITY CELL> that consist of two metal electrodes, usually stainless steel or titanium, in contact with the electrolyte solution. | [5 empty] | dm:ClassOfInanimatePhysicalObject
ANALYSIS MEASURING ELEMENT |  |  |  | CORROSION MEASURING ELEMENT | [5 empty] | rdl:RDS2225312 |  | A <CORROSION MEASURING ELEMENT> is an <ANALYSIS MEASURING ELEMENT> for the measurement of <CORROSION RATE>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | CORROSION MEASURING ELEMENT |  |  | CORROSION PROBE | [5 empty] | rdl:RDS648179 |  | A <CORROSION PROBE> is an <ANALYSIS MEASURING ELEMENT> that is a piece of suitable metal meant to corrode and give an indication of the corrosion rate. | [5 empty] | dm:ClassOfInanimatePhysicalObject
ENGINE COMPONENT |  |  |  | KNOCK SENSOR | [5 empty] | rdl:RDS6466164 |  | A <KNOCK SENSOR> is an <ENGINE COMPONENT> that is intended to detect engine knock in a combustion engine. | [5 empty] | dm:ClassOfInanimatePhysicalObject
ANALYSIS MEASURING ELEMENT |  |  |  | MOISTURE SENSING ELEMENT | [5 empty] | rdl:RDS5772924 |  | A <MOISTURE SENSING ELEMENT> is an <ANALYSIS MEASURING ELEMENT> for measurement of <MOISTURE> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | MOISTURE SENSING ELEMENT |  |  |  | CAPACITANCE MOISTURE SENSOR | [5 empty] | rdl:RDS11468322 |  | A <CAPACITANCE MOISTURE SENSOR> is an <INSTRUMENTATION ITEM> and a <MOISTURE SENSING ELEMENT> intended to determine moisture based on the capacitance principle. | [5 empty] | dm:ClassOfInanimatePhysicalObject
ANALYSIS MEASURING ELEMENT |  |  |  | PH MEASURING ELEMENT | [5 empty] | rdl:RDS2225313 |  | A <PH MEASURING ELEMENT> is an <ANALYSIS MEASURING ELEMENT> for the measurement of <PH>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | PH MEASURING ELEMENT |  |  | PH ELECTRODE | [5 empty] | rdl:RDS222505 |  | A <PH ELECTRODE> is an that is an electrode that uses a specially formulated, pH sensitive glass in contact with the solution, that develops a potential (voltage) proportional to the pH of the solution. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | DEW POINT MEASURING ELEMENT | [5 empty] | rdl:RDS2224465 |  | A <DEW POINT MEASURING ELEMENT> is a <MEASURING ELEMENT> for measuring of a <DEW POINT> | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | DISPLACEMENT MEASURING ELEMENT | [5 empty] | rdl:RDS5480833840 |  | A <DISPLACEMENT MEASURING ELEMENT> is a <MEASURING ELEMENT> for measurement of <DISPLACEMENT> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | DISPLACEMENT MEASURING ELEMENT |  |  | AXIAL DISPLACEMENT PROBE | [5 empty] | rdl:RDS6500874 |  | An <AXIAL DISPLACEMENT PROBE> is an <INSTRUMENTATION ITEM> and a <DISPLACEMENT MEASURING ELEMENT> for measurement of <AXIAL DISPLACEMENT>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | DISPLACEMENT MEASURING ELEMENT |  |  | RADIAL DISPLACEMENT PROBE | [5 empty] | rdl:RDS6500499 |  | An <RADIAL DISPLACEMENT PROBE> is an <INSTRUMENTATION ITEM> and a <DISPLACEMENT MEASURING ELEMENT> for measurement of <RADIAL DISPLACEMENT>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | DUST MEASURING ELEMENT | [5 empty] | rdl:RDS1299047481 |  | A <DUST MEASURING ELEMENT> is a <MEASURING ELEMENT> for measuring the velocity, mass, charge, flight direction, and number of dust particles, smaller than 100 micrometre, in space. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | DUST MEASURING ELEMENT |  |  | DUST MEASURING DEVICE | [5 empty] | rdl:RDS2225314 |  | A <DUST MEASURING DEVICE> is an <INSTRUMENTATION ITEM> and a <DUST MEASURING ELEMENT>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
DUST MEASURING ELEMENT |  |  |  | PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225315 |  | A <PARTICULATE MATTER MEASURING ELEMENT> is a <DUST MEASURING ELEMENT> for particles smaller than 10 micrometre, in space. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | OPTICAL PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225316 |  | An <OPTICAL PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for particle concentration that is based on an optical technology, such as light scatter. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | GRAVIMETRIC PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225317 |  | A <GRAVIMETRIC PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for particle concentratio that is based on gravimetric analysis. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | MICROBALANCE PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225318 |  | A <MICROBALANCE PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for particle concentratio that is based on a microbalance method. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | MICROSCOPICAL PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225319 |  | A <MICROSCOPICAL PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for size distribution that is based on microscopy. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | IMPACTION PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225320 |  | An <IMPACTION PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for size distribution that is based on impaction. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | DIFFUSION PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225321 |  | A <DIFFUSION PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for size distribution that is based on diffusion. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PARTICULATE MATTER MEASURING ELEMENT |  |  | CHARGING PARTICULATE MATTER MEASURING ELEMENT | [5 empty] | rdl:RDS2225322 |  | A <CHARGING PARTICULATE MATTER MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PARTICULAR MATTER MEASURING ELEMENT> for size distribution that is based on charging the partciles. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | ELECTRIC VARIABLE MEASURING ELEMENT | [5 empty] | rdl:RDS222506 |  | An <ELECTRIC VARIABLE MEASURING ELEMENT> is a <MEASURING ELEMENT> for measurement of an <ELECTRIC VARIABLE>. | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | CURRENT MEASURING ELEMENT | [5 empty] | rdl:RDS222507 |  | A <CURRENT MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <CURRENT> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | EARTH LEAKAGE MEASURING ELEMENT | [5 empty] | rdl:RDS2229115 |  | An <EARTH LEAKAGE MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <EARTH LEAKAGE> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | FREQUENCY MEASURING ELEMENT | [5 empty] | rdl:RDS2222222 |  | A <FREQUENCY MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <FREQUENCY> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | INSULATION FAULT MEASURING ELEMENT | [5 empty] | rdl:RDS2229116 |  | An <INSULATION FAULT MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <INSULATION FAULT> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | POWER FACTOR MEASURING ELEMENT | [5 empty] | rdl:RDS222508 |  | A <POWER FACTOR MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <POWER FACTOR> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | POWER MEASURING ELEMENT | [5 empty] | rdl:RDS1337135221 |  | A <POWER MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <POWER> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | RESISTANCE MEASURING ELEMENT | [5 empty] | rdl:RDS2222223 |  | A <RESISTANCE MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <RESISTANCE> | [5 empty] | dm:ClassOfFunctionalObject
ELECTRIC VARIABLE MEASURING ELEMENT |  |  |  | VOLTAGE MEASURING ELEMENT | [5 empty] | rdl:RDS1287147001 |  | A <VOLTAGE MEASURING ELEMENT> is an <ELECTRIC VARIABLE MEASURING ELEMENT> for measurement of <VOLTAGE> | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | EROSION DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS16757813 |  | An <EROSION DETECTING MEASURING ELEMENT> is a <MEASURING ELEMENT> that is designed to detect <EROSION>, e.g. caused by particles in a fluid | [5 empty] | dm:ClassOfFunctionalObject
EROSION DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | EROSION PROBE | [5 empty] | rdl:RDS5781161 |  | An <EROSION PROBE> is an <INSTRUMENTATION ITEM> and an <EROSION DETECTING MEASURING ELEMENT> that is detecting accumulated erosion directly | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | FIRE DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS1008629 |  | A <FIRE DETECTING MEASURING ELEMENT> is a <MEASURING ELEMENT> for detecting a <FIRE>, indirectly by means of detection of heat, smoke, flame or (fire) gas. | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2220149 |  | A <FLAME DETECTING MEASURING ELEMENT> is a <MEASURING ELEMENT> for detecting a <FLAME> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | ULTRAVIOLET FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225323 |  | An <ULTRAVIOLET FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting UV <RADIATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | NEAR IR ARRAY FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225324 |  | A <NEAR IR ARRAY FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting near infrared <RADIATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | INFRARED FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225325 |  | An <INFRARED FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting infrared <RADIATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | UV/IR FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225326 |  | An <UV/IR FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting both UV and IR <RADIATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | IR/IR FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225327 |  | An <IR/IR FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting IR <RADIATION> and comparing this with a reference frequency. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | IR3 FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225328 |  | An <IR3 FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting IR <RADIATION> and comparing this with a reference frequency above and below that IR <RADIATION>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLAME DETECTING MEASURING ELEMENT |  |  | 3IR+UV FLAME DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS2225329 |  | A <3IR+O42:O48FLAME DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLAME DETECTING MEASURING ELEMENT> that is based on detecting IR<RADIATION> with black body radiation compensation.. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | FLOW MEASURING ELEMENT | [5 empty] | rdl:RDS461834 |  | A <FLOW MEASURING ELEMENT> is a <MEASURING ELEMENT> that is used to measure <FLOW RATE> | [5 empty] | dm:ClassOfFunctionalObject
FLOW MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | MASS FLOW MEASURING ELEMENT | [5 empty] | rdl:RDS222509 |  | A <MASS FLOW MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLOW MEASURING ELEMENT> that is used to measure <MASS FLOW RATE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
MASS FLOW MEASURING ELEMENT |  |  |  |  | MASS FLOW MEASURING ELEMENT - CORIOLIS | [5 empty] | rdl:RDS613394 |  | A <MASS FLOW MEASURING ELEMENT - CORIOLIS> is a <MASS FLOW MEASURING ELEMENT> that uses the Coriolis effect that causes a laterally vibrating tube to distort, giving a direct measurement of mass flow. Furthermore, a direct measure of the density of the fluid is obtained | [5 empty] | dm:ClassOfInanimatePhysicalObject
MASS FLOW MEASURING ELEMENT |  |  |  |  | MASS FLOW MEASURING ELEMENT - THERMAL | [5 empty] | rdl:RDS883304 |  | A <MASS FLOW MEASURING ELEMENT - THERMAL> is a <MASS FLOW MEASURING ELEMENT> that uses combinations of heated elements and temperature sensors to measure the difference between static and flowing heat transfer to a fluid and infer its flow with a knowledge of the fluids specific heat and density. The fluid temperature is also measured and compensated for. If the density and specific heat characteristics of the fluid are constant, the meter can provide a direct mass flow readout, and does not need any additional pressure temperature compensation over their specified range | [5 empty] | dm:ClassOfInanimatePhysicalObject
FLOW MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | MULTIPHASE METER | [5 empty] | rdl:RDS1318334201 |  | A <MULTIPHASE METER> is an <INSTRUMENTATION ITEM> and a <FLOW MEASURING ELEMENT> for multiphase <METERING> of <LIQUID> and <GAS> flows | [5 empty] | dm:ClassOfInanimatePhysicalObject
FLOW MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | ORIFICE PLATE FLOWMETER ASSEMBLY | [5 empty] | rdl:RDS881909 |  | An <ORIFICE PLATE FLOWMETER ASSEMBLY> is an <INSTRUMENTATION ITEM> and a <FLOW MEASURING ELEMENT> that is an assembly of an orifice plate, carrier rings, flanges ans manifold as its main components | [5 empty] | dm:ClassOfInanimatePhysicalObject
FLOW MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | VOLUME FLOW MEASURING ELEMENT | [5 empty] | rdl:RDS222510 |  | A <VOLUME FLOW MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <FLOW MEASURING ELEMENT> that is used to measure <VOLUME FLOW RATE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | rdl:RDS418094 |  | A <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> is a <VOLUME FLOW MEASURING ELEMENT> that measures the volumetric flow rate of a liquid or gas by separating the flow stream into known volumes and counting them over time | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | VOLUME FLOW MEASURING ELEMENT - GEAR | [5 empty] | rdl:RDS222528 |  | A <VOLUME FLOW MEASURING ELEMENT - GEAR> is a <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> that uses gears to divide and meter the flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - GEAR | [6 empty] | VOLUME FLOW MEASURING ELEMENT - HELICAL GEAR | [5 empty] | rdl:RDS222529 |  | A <VOLUME FLOW MEASURING ELEMENT - HELICAL GEAR> is a <VOLUME FLOW MEASURING ELEMENT - GEAR> that is a flow meter that gets its name from the shape of its gears or rotors. These rotors resemble the shape of a helix, which is a spiral-shaped structure. As the fluid flows through the meter, it enters the compartments in the rotors, causing the rotors to rotate. Flowrate is calculated from the speed of rotation | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - GEAR | [6 empty] | VOLUME FLOW MEASURING ELEMENT - OVAL GEAR | [5 empty] | rdl:RDS880424 |  | A <VOLUME FLOW MEASURING ELEMENT - OVAL GEAR> is a <VOLUME FLOW MEASURING ELEMENT - GEAR> that is a positive displacement meter that uses two or more oblong gears configured to rotate at right angles to one another, forming a T shape | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | VOLUME FLOW MEASURING ELEMENT - HELIX | [5 empty] | rdl:RDS222530 |  | A <VOLUME FLOW MEASURING ELEMENT - HELIX> is a <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> that is a type of turbine flow meter that comprises a rotor with helical blades inserted axially in the flow, much like a ducted fan | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ROTARY | [5 empty] | rdl:RDS222531 |  | A <VOLUME FLOW MEASURING ELEMENT - ROTARY> is a <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> that uses eccentrically mounted drums to divide and meter the flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | VOLUME FLOW MEASURING ELEMENT - NUTATING DISC | [5 empty] | rdl:RDS881819 |  | A <VOLUME FLOW MEASURING ELEMENT - NUTATING DISC> is a <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> that uses a disk which fits closely in a measuring chamber. As the fluid enters one side of the partition and leaves through the other side, it makes the disk to wobble. The number of revolutions made by a small shaft attached to the disk will indicate the volume of flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT | [5 empty] | VOLUME FLOW MEASURING ELEMENT - PISTON | [5 empty] | rdl:RDS812114 |  | A <VOLUME FLOW MEASURING ELEMENT - PISTON> is a <VOLUME FLOW MEASURING ELEMENT - POSITIVE DISPLACEMENT> that is utilizing a piston closely fitted into a cylindrical chamber with a pair of slots that are progressively uncovered as the pressure differential makes the piston oscillate/spin around a control roller in the centre of the chamber | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE | [5 empty] | rdl:RDS222516 |  | A <VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE> is a <VOLUME FLOW MEASURING ELEMENT> that measures the force exerted by a fluid in order to calculate its flow rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - IMPACT | [5 empty] | rdl:RDS222517 |  | A <VOLUME FLOW MEASURING ELEMENT - IMPACT> is a <VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE> that measures the impact force of a fluid on a plate and outputs a signal representative of flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - TARGET | [5 empty] | rdl:RDS222518 |  | A <VOLUME FLOW MEASURING ELEMENT - TARGET> is a <VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE> that uses a target, with force bar, as MEASURING ELEMENT | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - VANE | [5 empty] | rdl:RDS222519 |  | A <VOLUME FLOW MEASURING ELEMENT - VANE> is a <VOLUME FLOW MEASURING ELEMENT - FORCE-BALANCE> that uses one or more vanes in a measuring chamber to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA | [5 empty] | rdl:RDS222515 |  | A <VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA> is a <VOLUME FLOW MEASURING ELEMENT> that measures fluid flow by allowing the cross sectional area of the device to vary in response to the flow, causing some measurable effect that indicates the rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA | [5 empty] | FLOAT COLLAR | [5 empty] | rdl:RDS5786112 |  | A <FLOAT COLLAR> is a <VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA> that is a specially formed float for use in variable area flow meters and variable area flow indicators that has a collar or flange on the top | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ROTAMETER | [5 empty] | rdl:RDS882719 |  | A <VOLUME FLOW MEASURING ELEMENT - ROTAMETER> is a <VOLUME FLOW MEASURING ELEMENT - VARIABLE AREA> that consists of a tapered tube, typically made of glass, with a float inside that is pushed up by fluid flow and pulled down by gravity. As flow rate increases, greater viscous and pressure forces on the float cause it to rise until it becomes stationary at a location in the tube that is wide enough for the forces to balance | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | rdl:RDS222532 |  | A <VOLUME FLOW MEASURING ELEMENT - VELOCITY> is a <VOLUME FLOW MEASURING ELEMENT> that measures the velocity of a fluid in order to calculate its flow rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ACOUSTIC DOPPLER | [5 empty] | rdl:RDS222533 |  | A <VOLUME FLOW MEASURING ELEMENT - ACOUSTIC DOPPLER> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> measuring the velocity of particles in a remote sampling volume based upon the Doppler shift effect; records velocity components at a single point with a relatively high frequency | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ELECTROMAGNETIC | [5 empty] | rdl:RDS1009664 |  | A <VOLUME FLOW MEASURING ELEMENT - ELECTROMAGNETIC> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that uses a magnetic field applied to the metering tube, which results in a potential difference proportional to the flow velocity perpendicular to the flux lines. The potential difference is sensed by electrodes aligned perpendicular to the flow and the applied magnetic field. The physical principle at work is Faraday&amp;#39;s law of electromagnetic induction. The magnetic flow meter requires a conducting fluid and a nonconducting pipe liner | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - LASER DOPPLER | [5 empty] | rdl:RDS222535 |  | A <VOLUME FLOW MEASURING ELEMENT - LASER DOPPLER> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that uses a beam of laser light impinging on a moving particle; the light will be partially scattered with a change in wavelength proportional to the particles speed (the Doppler effect) | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - MULTIPLE JET | [5 empty] | rdl:RDS222536 |  | A <VOLUME FLOW MEASURING ELEMENT - MULTIPLE JET> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that is a velocity type meter that has an impeller that rotates horizontally on a vertical shaft. The impeller element is in a housing in which multiple inlet ports direct the fluid flow at the impeller causing it to rotate in a specific direction in proportion to the flow velocity. This meter works mechanically much like a single jet meter except that the ports direct the flow at the impeller equally from several points around the circumference of the element, not just one point; this minimizes uneven wear on the impeller and its shaft | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - OPTICAL | [5 empty] | rdl:RDS222537 |  | A <VOLUME FLOW MEASURING ELEMENT - OPTICAL> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that is a flow meter that uses light to determine flow rate. Small particles which accompany natural and industrial gases pass through two laser beams focused a short distance apart in the flow path. Laser light is scattered when a particle crosses the first beam. The detecting optics collects scattered light on a photodetector, which then generates a pulse signal. As the same particle crosses the second beam, the detecting optics collect scattered light on a second photodetector, which converts the incoming light into a second electrical pulse. By measuring the time interval between these pulses, the gas velocity is calculated | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [5 empty] | rdl:RDS222540 |  | A <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that uses the rotation of a shaped element to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - PADDLE INSERTION | [5 empty] | rdl:RDS222541 |  | A <VOLUME FLOW MEASURING ELEMENT - PADDLE INSERTION> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> that uses the rotation of a paddle wheel to measure flow at a point inside the process | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - PROPELLER | [5 empty] | rdl:RDS222542 |  | A <VOLUME FLOW MEASURING ELEMENT - PROPELLER> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> that uses the rotation of a propeller to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - RADIAL TURBINE | [5 empty] | rdl:RDS222543 |  | A <VOLUME FLOW MEASURING ELEMENT - RADIAL TURBINE> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> where a Pelton wheel rotates in the liquid flow around an axis into a user-readable rate of flow. The Pelton wheel tends to have all the flow traveling around it with the inlet flow focused on the blades by a jet | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - ROTATING VANE | [5 empty] | rdl:RDS222544 |  | A <VOLUME FLOW MEASURING ELEMENT - ROTATING VANE> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> that uses the rotation of a vane to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - TURBINE | [5 empty] | rdl:RDS417914 |  | A <VOLUME FLOW MEASURING ELEMENT - TURBINE> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> that translates the mechanical action of the turbine rotating in the liquid flow around an axis into a user-readable rate of flow (gpm, lpm, etc.). The turbine tends to have all the flow traveling around it. The turbine wheel is set in the path of a fluid stream | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT | [6 empty] | VOLUME FLOW MEASURING ELEMENT - PADDLE WHEEL | [5 empty] | rdl:RDS882809 |  | A <VOLUME FLOW MEASURING ELEMENT - PADDLE WHEEL> is a <VOLUME FLOW MEASURING ELEMENT - ROTATING ELEMENT> that consists of an impeller is small with respect to the width of the pipe, and projects only partially into the flow, like the paddle wheel | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - SINGLE JET | [5 empty] | rdl:RDS222539 |  | A <VOLUME FLOW MEASURING ELEMENT - SINGLE JET> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that consists of an impeller with radial vanes, impinged upon by a single jet | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - SWIRL | [5 empty] | rdl:RDS222545 |  | A <VOLUME FLOW MEASURING ELEMENT - SWIRL> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that uses swirls generated by fixed spiral vanes to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - VORTEX | [5 empty] | rdl:RDS418184 |  | A <VOLUME FLOW MEASURING ELEMENT - VORTEX> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that involves placing a bluff body (called a shedder bar) in the path of the fluid. As the fluid passes this bar, disturbances in the flow called vortices are created. The vortices trail behind the cylinder, alternatively from each side of the bluff body. The frequency at which these vortices alternate sides is essentially proportional to the flow rate of the fluid. Inside, atop, or downstream of the shedder bar is a sensor for measuring the frequency of the vortex shedding. This sensor is often a piezoelectric crystal, which produces a small, but measurable, voltage pulse every time a vortex is created. Since the frequency of such a voltage pulse is also proportional to the fluid velocity, a volumetric flow rate is calculated using the cross sectional area of the flow meter | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ELECTROMAGNETIC INSERTION | [5 empty] | rdl:RDS222534 |  | A <VOLUME FLOW MEASURING ELEMENT - ELECTROMAGNETIC INSERTION> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that uses electromagnetic principles to measure flow at a point inside the process | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ANEMOMETRIC | [5 empty] | rdl:RDS5773048 |  | A <VOLUME FLOW MEASURING ELEMENT - ANEMOMETRIC> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that can be the sensing element of an ANEMOMETER or wind gauge | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - VELOCITY | [5 empty] | VOLUME FLOW MEASURING ELEMENT - ULTRASONIC | [5 empty] | rdl:RDS2224558 |  | A <VOLUME FLOW MEASURING ELEMENT - ULTRASONIC> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> based on ultrasonic technology | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ULTRASONIC | [6 empty] | VOLUME FLOW MEASURING ELEMENT - ULTRASONIC DOPPLER | [5 empty] | rdl:RDS1009799 |  | A <VOLUME FLOW MEASURING ELEMENT - ULTRASONIC DOPPLER> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that measures the Doppler shift resulting from reflecting an ultrasonic beam off the particulates in flowing fluid. The frequency of the transmitted beam is affected by the movement of the particles; this frequency shift can be used to calculate the fluid velocity. For the Doppler principle to work there must be a high enough density of sonically reflective materials such as solid particles or air bubbles suspended in the fluid | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ULTRASONIC | [6 empty] | VOLUME FLOW MEASURING ELEMENT - ULTRASONIC TRANSIT TIME | [5 empty] | rdl:RDS418139 |  | A <VOLUME FLOW MEASURING ELEMENT - ULTRASONIC TRANSIT TIME> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that measures the difference of the transit time of ultrasonic pulses propagating in and against the direction of flow. This time difference is a measure for the average velocity of the fluid along the path of the ultrasonic beam | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ULTRASONIC TRANSIT TIME | [7 empty] | TIME OF FLIGHT ULTRASONIC FLOW METER | [5 empty] | rdl:RDS883349 |  | A <TIME OF FLIGHT ULTRASONIC FLOW METER> is an <VOLUME FLOW MEASURING ELEMENT - ULTRASONIC TRANSIT TIME> that use the time for the sound to travel between a transmitter and a receiver to measure a flow rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - ULTRASONIC | [6 empty] | VOLUME FLOW MEASURING ELEMENT - ULTRASONIC IN-LINE | [5 empty] | rdl:RDS880559 |  | A <VOLUME FLOW MEASURING ELEMENT - ULTRASONIC IN-LINE> is a <VOLUME FLOW MEASURING ELEMENT - VELOCITY> that is mounted in-line with a pipeline | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL | [5 empty] | rdl:RDS222520 |  | A <VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL> is a <VOLUME FLOW MEASURING ELEMENT< where behind a <WEIR> or in a <FLUME> the level of the water is measured at a designated point in a hydraulic structure using various secondary devices (bubblers, ultrasonic, float, and differential pressure are common methods). This level is converted to a flow rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL | [5 empty] | VOLUME FLOW MEASURING ELEMENT - CURRENT | [5 empty] | rdl:RDS222521 |  | A <VOLUME FLOW MEASURING ELEMENT - CURRENT> is a <VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN> CHANNEL that where a propeller is traversed over the area of a channel and velocities averaged to calculate total flow. This may be on the order of hundreds of cubic meters per second. The flow must be kept steady during the traverse of the current meters. Methods for testing hydroelectric turbines are given in IEC standard 60041 | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL | [5 empty] | VOLUME FLOW MEASURING ELEMENT - FLUME | [5 empty] | rdl:RDS222522 |  | A <VOLUME FLOW MEASURING ELEMENT - FLUME> is a <VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN> CHANNEL that is a specially shaped, fixed hydraulic structure that under free-flow conditions forces flow to accelerate in such a manner that the flow rate through the flume can be characterized by a level-to-flow relationship as applied to a single head (level) measurement within the flume. Acceleration is accomplished through a convergence of the sidewalls, a change in floor elevation, or a combination of the two | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL | [5 empty] | VOLUME FLOW MEASURING ELEMENT - WEIR | [5 empty] | rdl:RDS222524 |  | A <VOLUME FLOW MEASURING ELEMENT - WEIR> is a <VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN> CHANNEL that is an obstruction in the flow path in an open channel. The weir will cause an increase in the water depth as the water flows over the weir. In general, the greater the flow rate, the greater will be the increase in depth of flow, The height of water above the top of the weir is the measurement usually used to correlate with flow rate | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - WEIR | [6 empty] | VOLUME FLOW MEASURING ELEMENT - BROAD-CRESTED WEIR | [5 empty] | rdl:RDS222525 |  | A <VOLUME FLOW MEASURING ELEMENT - BROAD-CRESTED WEIR> is a <VOLUME FLOW MEASURING ELEMENT - WEIR> that is a flat-crested structure, with a long crest compared to the flow thickness | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - WEIR | [6 empty] | VOLUME FLOW MEASURING ELEMENT - SHARP-CRESTED WEIR | [5 empty] | rdl:RDS222526 |  | A <VOLUME FLOW MEASURING ELEMENT - SHARP-CRESTED WEIR> is a <VOLUME FLOW MEASURING ELEMENT - WEIR> that is allowing the fluid to fall cleanly away from the weir. Sharp crested weirs are typically 1/4 inch or thinner metal plates | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - WEIR | [6 empty] | VOLUME FLOW MEASURING ELEMENT - V-NOTCH WEIR | [5 empty] | rdl:RDS222527 |  | A <VOLUME FLOW MEASURING ELEMENT - V-NOTCH WEIR> is a <VOLUME FLOW MEASURING ELEMENT - WEIR> that is a triangular channel section, used to measure small discharge values. The upper edge of the section is always above the water level, so the channel is always triangular simplifying calculation of the cross-sectional area | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN CHANNEL | [5 empty] | VOLUME FLOW MEASURING ELEMENT - AREA-VELOCITY | [5 empty] | rdl:RDS881864 |  | A <VOLUME FLOW MEASURING ELEMENT - AREA-VELOCITY> is a <VOLUME FLOW MEASURING ELEMENT - LEVEL-TO-FLOW OPEN> CHANNEL that calculates flow in an open channel by measuring the level to determine the cross section, and it measures the fluid velocity; subequently these two values are multiplied to arrive at the <VOLUME FLOW RATE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - AREA-VELOCITY | [6 empty] | WEIR FLOW METER | [5 empty] | rdl:RDS883754 |  | A <WEIR FLOW METER> is an <OPEN CHANNEL FLOW METER> utilizing some kind of weir as obstruction | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | rdl:RDS222511 |  | A <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> is a <VOLUME FLOW MEASURING ELEMENT> that measures the differential pressure across a MEASURING ELEMENT in order to measure flow | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | ORIFICE PLATE | [5 empty] | rdl:RDS418364 |  | An <ORIFICE PLATE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that is a plate with a hole through it, placed in the flow; it constricts the flow, and measuring the pressure differential across the constriction gives the flow rate. It is basically a crude form of Venturi meter, but with higher energy losses. There are three type of orifice: concentric, eccentric, and segmental | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | CONCENTRIC HOLE ORIFICE PLATE | [5 empty] | rdl:RDS7199620 |  | A <CONCENTRIC HOLE ORIFICE PLATE> is an <ORIFICE PLATE> with a circular orifice that is coaxial with the conduit and is concentric to the conduit axis | [5 empty] | dm:ClassOfInanimatePhysicalObject
CONCENTRIC HOLE ORIFICE PLATE | [7 empty] | CONICAL CONCENTRIC ORIFICE PLATE | [5 empty] | rdl:RDS7199744 |  | A <CONICAL CONCENTRIC ORIFICE PLATE> is a <CONCENTRIC HOLE ORIFICE PLATE> of which the orifice is conical | [5 empty] | dm:ClassOfInanimatePhysicalObject
CONCENTRIC HOLE ORIFICE PLATE | [7 empty] | CYLINDRICAL CONCENTRIC ORIFICE PLATE | [5 empty] | rdl:RDS7200382 |  | A <CYLINDRICAL CONCENTRIC ORIFICE PLATE> is a <CONCENTRIC HOLE ORIFICE PLATE> of which the orifice is cylindrical | [5 empty] | dm:ClassOfInanimatePhysicalObject
CONCENTRIC HOLE ORIFICE PLATE | [7 empty] | QUARTER-CIRCLE CONCENTRIC ORIFICE PLATE | [5 empty] | rdl:RDS7200320 |  | A <QUARTER-CIRCLE CONCENTRIC ORIFICE PLATE> is a <CONCENTRIC HOLE ORIFICE PLATE> of which the orifice is a quarter of the circumference of a circle | [5 empty] | dm:ClassOfInanimatePhysicalObject
CONCENTRIC HOLE ORIFICE PLATE | [7 empty] | SQUARE-EDGED CONCENTRIC ORIFICE PLATE | [5 empty] | rdl:RDS7199682 |  | A <SQUARE-EDGED CONCENTRIC ORIFICE PLATE> is a <CONCENTRIC HOLE ORIFICE PLATE> of which the orifice is square-shaped | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | ECCENTRIC HOLE ORIFICE PLATE | [5 empty] | rdl:RDS7199806 |  | An <ECCENTRIC HOLE ORIFICE PLATE> is an <ORIFICE PLATE> with an orifice that is circular and coaxial with the conduit, but eccentric to the conduit axis | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | PADDLE ORIFICE PLATE | [5 empty] | rdl:RDS6759480 |  | A <PADDLE ORIFICE PLATE> is an <ORIFICE PLATE> with an integrated paddle or handle to facilitate installation between flanges | [5 empty] | dm:ClassOfInanimatePhysicalObject
PADDLE ORIFICE PLATE | [7 empty] | INTEGRAL PADDLE ORIFICE PLATE | [5 empty] | rdl:RDS7211344 |  | An <INTEGRAL PADDLE ORIFICE PLATE> is a <PADDLE ORIFICE PLATE> that is made from one piece | [5 empty] | dm:ClassOfInanimatePhysicalObject
PADDLE ORIFICE PLATE | [7 empty] | WELDED PADDLE ORIFICE PLATE | [5 empty] | rdl:RDS7211282 |  | A <WELDED PADDLE ORIFICE PLATE> is a <PADDLE ORIFICE PLATE> where the orifce plate and the paddle (handle) is assembled through welding | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | QUARTER CIRCLE ORIFICE PLATE | [5 empty] | rdl:RDS7022320 |  | A <QUARTER CIRCLE ORIFICE PLATE> is an <ORIFICE PLATE> with an orifice that has the shape of a quarter circle | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | RESTRICTION ORIFICE | [5 empty] | rdl:RDS882629 |  | A <RESTRICTION ORIFICE> is an <ORIFICE PLATE> that is intended for use as a restrictor | [5 empty] | dm:ClassOfInanimatePhysicalObject
RESTRICTION ORIFICE | [7 empty] | MULTI-HOLE RESTRICTION ORIFICE | [5 empty] | rdl:RDS1465077151 |  | A <MULTI-HOLE RESTRICTION ORIFICE> is a <RESTRICTION ORIFICE> with more than one hole | [5 empty] | dm:ClassOfInanimatePhysicalObject
RESTRICTION ORIFICE | [7 empty] | SINGLE-HOLE RESTRICTION ORIFICE | [5 empty] | rdl:RDS1465078021 |  | A <SINGLE-HOLE RESTRICTION ORIFICE> is a <RESTRICTION ORIFICE PLATE> with one hole | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | SEAL ORIFICE | [5 empty] | rdl:RDS7042877 |  | A <SEAL ORIFICE> is an <ORIFICE PLATE> intended to limit the flow of seal flush to the seal | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | SEGMENTAL ORIFICE PLATE | [5 empty] | rdl:RDS7199868 |  | A <SEGMENTAL ORIFICE PLATE> is an <ORIFICE PLATE> with an orifice that has the form of a segment of a circle | [5 empty] | dm:ClassOfInanimatePhysicalObject
SEGMENTAL ORIFICE PLATE | [7 empty] | SEGMENTAL ORIFICE PLATE FOR DIRTY GAS | [5 empty] | rdl:RDS7199960 |  | A <SEGMENTAL ORIFICE PLATE FOR DIRTY GAS> is a <SEGMENTAL ORIFICE PLATE> that is in the form of a segment of a circle, the cord of which is vertical. The hole is tangential to the side of the conduit | [5 empty] | dm:ClassOfInanimatePhysicalObject
SEGMENTAL ORIFICE PLATE | [7 empty] | SEGMENTAL ORIFICE PLATE FOR DIRTY LIQUID | [5 empty] | rdl:RDS7200022 |  | A <SEGMENTAL ORIFICE PLATE FOR DIRTY LIQUID> is a <SEGMENTAL ORIFICE PLATE> that is in the form of a segment of a circle, the cord of which is horizontal. The hole is tangential to the bottom of the conduit | [5 empty] | dm:ClassOfInanimatePhysicalObject
ORIFICE PLATE | [6 empty] | UNIVERSAL TYPE ORIFICE PLATE | [5 empty] | rdl:RDS6759638 |  | A <UNIVERSAL TYPE ORIFICE PLATE> is an <ORIFICE PLATE> with a circular design that is fit for use in orifice fittings or in the plate holders of ring-type joint (RTJ) orifice flanges | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VENTURI TUBE ASSEMBLY | [5 empty] | rdl:RDS7940784 |  | A <VENTURI TUBE ASSEMBLY> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that is an assembly of a venturi tube, flow straightener and meter tube as its main components | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - CONE | [5 empty] | rdl:RDS222512 |  | A <VOLUME FLOW MEASURING ELEMENT - CONE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that works with the same basic principles as Venturi and Orifice type DP meters, but don&amp;#226;&#226;â€š&#172;&#226;â€ž&#162;t require the same upstream and downstream piping. The cone acts as a conditioning device as well as a differential pressure producer | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - PITOT TUBE | [5 empty] | rdl:RDS418274 |  | A <VOLUME FLOW MEASURING ELEMENT - PITOT TUBE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that is a pressure measuring device used to measure fluid flow velocity by determining the stagnation pressure and static pressure. Bernoulli's equation used to calculate the dynamic pressure and hence fluid velocity | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - PITOT TUBE | [6 empty] | MULTIPORT PITOT TUBE | [5 empty] | rdl:RDS418319 |  | A <MULTIPORT PITOT TUBE> is a <VOLUME FLOW MEASURING ELEMENT - PITOT TUBE> that has multiple impact pressure ports normal to the flow direction on the exterior tube | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - SEGMENTAL WEDGE | [5 empty] | rdl:RDS222513 |  | A <VOLUME FLOW MEASURING ELEMENT - SEGMENTAL WEDGE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that consists of a cylindrical pressure vessel into which a constriction is fabricated thereby leaving an open segment of a known height. Pressure taps on either side of the wedge provide the differential signal | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - V-CONE | [5 empty] | rdl:RDS222514 |  | A <VOLUME FLOW MEASURING ELEMENT - V-CONE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that uses a conical body as MEASURING ELEMENT | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - DALL TUBE | [5 empty] | rdl:RDS7023655 |  | A <VOLUME FLOW MEASURING ELEMENT - DALL TUBE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that is a shortened version of a Venturi meter, with a lower pressure drop than an orifice plate. As with these flow meters the flow rate in a Dall tube is determined by measuring the pressure drop caused by restriction in the conduit | [5 empty] | dm:ClassOfInanimatePhysicalObject
VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | VOLUME FLOW MEASURING ELEMENT - VENTURI TUBE | [5 empty] | rdl:RDS648044 |  | A <VOLUME FLOW MEASURING ELEMENT - VENTURI TUBE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> that constricts the flow in some fashion, and pressure sensors measure the differential pressure before and within the constriction. This method is widely used to measure flow rate in the transmission of gas through pipelines | [5 empty] | dm:ClassOfInanimatePhysicalObject
NOZZLE(fluid velocity) | VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE |  |  |  |  | VOLUME FLOW MEASURING ELEMENT - FLOW NOZZLE | [5 empty] | rdl:RDS2223661 |  | A <VOLUME FLOW MEASURING ELEMENT - FLOW NOZZLE> is a <VOLUME FLOW MEASURING ELEMENT - DIFFERENTIAL PRESSURE> and a <NOZZLE(fluid velocity)> that consists of a convergent inlet whose shape is a quarter ellipse, and a cylindrical throat | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | FLUID LEAK DETECTION MEASURING ELEMENT | [5 empty] | rdl:RDS222552 |  | A <FLUID LEAK DETECTION MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of a fluid leak | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | FLUID LEAK DETECTION MEASURING ELEMENT |  |  | GAS LEAK DETECTION SENSOR | [5 empty] | rdl:RDS2225330 |  | A <GAS LEAK DETECTION SENSOR> is an <INSTRUMENTATION ITEM> and a <FLUID LEAK DETECTION MEASURING ELEMENT> that is sensing a <GAS> leakage. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLUID LEAK DETECTION MEASURING ELEMENT |  |  | LIQUID LEAK DETECTION SENSOR | [5 empty] | rdl:RDS2225331 |  | A <LIQUID LEAK DETECTION SENSOR> is an <INSTRUMENTATION ITEM> and a <FLUID LEAK DETECTION MEASURING ELEMENT> that is sensing a <LIQUID> leakage. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LIQUID LEAK DETECTION SENSOR |  |  |  | WATER DETECTOR | [5 empty] | rdl:RDS883709 |  | A <WATER DETECTOR> is an <INSTRUMENTATION ITEM> and a <MEASURING ELEMENT> that is designed to detect the presence of water and provide input to an alert to allow the prevention of water damage. A common design is a small cable or device that lies flat on a floor and relies on the electrical conductivity of water to decrease the resistance across two contacts | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FLUID LEAK DETECTION MEASURING ELEMENT |  |  | VAPOUR LEAK DETECTION SENSOR | [5 empty] | rdl:RDS2225332 |  | A <VAPOUR LEAK DETECTION SENSOR> is an <INSTRUMENTATION ITEM> and a <FLUID LEAK DETECTION MEASURING ELEMENT> that is sensing a <VAPOUR> leakage. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | FORCE MEASURING ELEMENT | [5 empty] | rdl:RDS1337126251 |  | A <FORCE MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <FORCE> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | FORCE MEASURING ELEMENT |  |  | DIFFERENTIAL FORCE SENSOR | [5 empty] | rdl:RDS1337128971 |  | A <DIFFERENTIAL FORCE SENSOR> is an <INSTRUMENTATION ITEM> and a <FORCE MEASURING ELEMENT> that is sensing a difference in force and produces a corresponding output | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | FORCE MEASURING ELEMENT |  |  | FORCE SENSOR | [5 empty] | rdl:RDS222546 |  | A <FORCE SENSOR> is an <INSTRUMENTATION ITEM> and a <FORCE MEASURING ELEMENT that is a piezoresistive force sensor | [5 empty] | dm:ClassOfInanimatePhysicalObject
FORCE SENSOR |  |  |  |  | STRAIN GAUGE LOAD CELL | [5 empty] | rdl:RDS5770253 |  | A <STRAIN GAUGE LOAD CELL> is a <FORCE SENSOR> that measures the deformation (strain) as a change in electrical resistance, which is a measure of the strain and hence the applied forces | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | GAS DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS477134 |  | A <GAS DETECTING MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of the presence of a specified concentration of a gas | [5 empty] | dm:ClassOfFunctionalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - ELECTROCHEMICAL | [5 empty] | rdl:RDS222547 |  | A <GAS DETECTING MEASURING ELEMENT - ELECTROCHEMICAL> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT> where gases are allowed to diffuse through a porous membrane to an electrode where it is either chemically oxidized or reduced. The amount of current produced is determined by how much of the gas is oxidized at the electrode,[3] indicating the concentration of the gas | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - HOLOGRAPHIC | [5 empty] | rdl:RDS222548 |  | A <GAS DETECTING MEASURING ELEMENT - HOLOGRAPHIC> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT>using light reflection to detect changes in a polymer film matrix containing a hologram. Since holograms reflect light at certain wavelengths, a change in their composition can generate a colorful reflection indicating the presence of a gas molecule | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - INFRARED IMAGING | [5 empty] | rdl:RDS222549 |  | A <GAS DETECTING MEASURING ELEMENT - INFRARED IMAGING> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT> that can be an active or passive system. For active sensing, IR imaging sensors typically scan a laser across the field of view of a scene and look for backscattered light at the absorption line wavelength of a specific target gas. Passive IR imaging sensors measure spectral changes at each pixel in an image and look for specific spectral signatures that indicate the presence of target gases | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - SEMICONDUCTOR | [5 empty] | rdl:RDS222550 |  | A <GAS DETECTING MEASURING ELEMENT - SEMICONDUCTOR> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT> using a chemical reaction that takes place, causing the electrical resistance in the sensor to be decreased when it comes in contact with the monitored gas | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - ULTRASONIC | [5 empty] | rdl:RDS222551 |  | A <GAS DETECTING MEASURING ELEMENT - ULTRASONIC> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT> using acoustic sensors to detect changes in the background noise of its environment | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAS DETECTING MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | GAS DETECTING MEASURING ELEMENT - INFRARED POINT | [5 empty] | rdl:RDS7477380 |  | A <GAS DETECTING MEASURING ELEMENT - INFRARED POINT> is an <INSTRUMENTATION ITEM> and a <GAS DETECTING MEASURING ELEMENT> using radiation passing through a known volume of gas; energy from the sensor beam is absorbed at certain wavelengths, depending on the properties of the specific gas | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | HEAT DETECTION MEASURING ELEMENT | [5 empty] | rdl:RDS2229121 |  | A <HEAT DETECTION MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of a rising heat | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | HUMIDITY MEASURING ELEMENT | [5 empty] | rdl:RDS2229123 |  | A <HUMIDITY MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of a <HUMIDITY> | [5 empty] | dm:ClassOfFunctionalObject
HUMIDITY MEASURING ELEMENT |  |  |  | ABSOLUTE HUMIDITY MEASURING ELEMENT | [5 empty] | rdl:RDS2225333 |  | An <ABSOLUTE HUMIDITY MEASURING ELEMENT> is a <HUMIDITY MEASURING ELEMENT> for the measurement of <ABSOLUTE HUMIDITY>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | ABSOLUTE HUMIDITY MEASURING ELEMENT |  |  | ABSOLUTE HUMIDITY PROBE | [5 empty] | rdl:RDS222504 |  | An <ABSOLUTE HUMIDITY PROBE> is an <INSTRUMENTATION ITEM> and an <ABSOLUTE HUMIDITY MEASURING ELEMENT> for measurement of <ABSOLUTE HUMIDITY>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
HUMIDITY MEASURING ELEMENT |  |  |  | RELATIVE HUMIDITY MEASURING ELEMENT | [5 empty] | rdl:RDS2225334 |  | A <RELATIVE HUMIDITY MEASURING ELEMENT> is a <HUMIDITY MEASURING ELEMENT> for the measurement of <RELATIVE HUMIDITY>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | RELATIVE HUMIDITY MEASURING ELEMENT |  |  | RELATIVE HUMITITY PROBE | [5 empty] | rdl:RDS5772862 |  | A <RELATIVE HUMIDITY PROBE> is an <INSTRUMENTATION ITEM> and an <RELATIVE HUMIDITY MEASURING ELEMENT> for measurement of <RELATIVE HUMIDITY>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT | [5 empty] | rdl:RDS880784 |  | A <LEVEL MEASURING ELEMENT> is a <MEASURING ELEMENT> for measuring a <LIQUID LEVEL>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - DISPLACER | [5 empty] | rdl:RDS222553 |  | A <LEVEL MEASURING ELEMENT - DISPLACER> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that uses the buoyancy force from a suspended displacer that is immersed in the liquid | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY | [5 empty] | rdl:RDS222554 |  | A <LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that uses the electric properties of a process material to measure its level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY |  |  |  |  | LEVEL MEASURING ELEMENT - CONDUCTANCE | [5 empty] | rdl:RDS222555 |  | A <LEVEL MEASURING ELEMENT - CONDUCTANCE> is a <LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY> that uses a low-voltage, current-limited power source applied across separate electrodes. The power supply is matched to the conductivity of the liquid, with higher voltage versions designed to operate in less conductive (higher resistance) mediums. A conductive liquid contacting both the longest probe (common) and a shorter probe (return) completes a conductive circuit | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY |  |  |  |  | LEVEL MEASURING ELEMENT - RF ADMITTANCE | [5 empty] | rdl:RDS222559 |  | A <LEVEL MEASURING ELEMENT - RF ADMITTANCE> is a <LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY> that uses a rod probe and RF source to measures the change in admittance (a measure of how easily a circuit will allow a current to flow). When the level changes around the probe, a corresponding change in the di-electric is observed. This changes the admittance of this imperfect capacitor and this change is measured to detect change of level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY |  |  |  |  | LEVEL MEASURING ELEMENT - CAPACITANCE | [5 empty] | rdl:RDS6500394 |  | A <LEVEL MEASURING ELEMENT - CAPACITANCE> is a <LEVEL MEASURING ELEMENT - ELECTRIC PROPERTY> that senses the presence of a wide variety of solids, aqueous and organic liquids, and slurries. The technique is frequently referred to as RF for the radio frequency signals applied to the capacitance circuit | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE | [5 empty] | rdl:RDS222563 |  | A <LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that is a submersible or externally mounted pressure sensor that senses increasing pressure with depth | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE |  |  |  |  | LEVEL MEASURING ELEMENT - DIFFERENTIAL PRESSURE | [5 empty] | rdl:RDS222564 |  | A <LEVEL MEASURING ELEMENT - DIFFERENTIAL PRESSURE> is a <LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE> that is measuring liquid level by sensing pressure difference between two points, e.g. the hydrostatic pressure at the bottom of a tank to the atmosphere pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE |  |  |  |  | LEVEL MEASURING ELEMENT - PRESSURE | [5 empty] | rdl:RDS222565 |  | A <LEVEL MEASURING ELEMENT - PRESSURE> is a <LEVEL MEASURING ELEMENT - HYDROSTATIC PRESSURE> that uses gauge pressure to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - PRESSURE | [5 empty] | LEVEL MEASURING ELEMENT - AIR BUBBLER SYSTEM | [5 empty] | rdl:RDS222566 |  | A <LEVEL MEASURING ELEMENT - AIR BUBBLER SYSTEM> is a <LEVEL MEASURING ELEMENT - PRESSURE> that uses a tube with an opening below the surface of the liquid level. A fixed flow of air is passed through the tube. Pressure in the tube is proportional to the depth (and density) of the liquid over the outlet of the tube | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - PRESSURE | [5 empty] | LEVEL MEASURING ELEMENT - IMMERSION | [5 empty] | rdl:RDS222567 |  | A <LEVEL MEASURING ELEMENT - IMMERSION> is a <LEVEL MEASURING ELEMENT - PRESSURE> the sensing element of which is immersed in the liquid to be measured | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - MECHANICAL | [5 empty] | rdl:RDS222568 |  | A <LEVEL MEASURING ELEMENT - MECHANICAL> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that uses mechanical means to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - MECHANICAL |  |  |  |  | LEVEL MEASURING ELEMENT - PLUMB BOB | [5 empty] | rdl:RDS222569 |  | A <LEVEL MEASURING ELEMENT - PLUMB BOB> is a <LEVEL MEASURING ELEMENT - MECHANICAL> that uses a wire or tape with a body as sensing element to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - MECHANICAL |  |  |  |  | LEVEL MEASURING ELEMENT - ROTATING PADDLE | [5 empty] | rdl:RDS222570 |  | A <LEVEL MEASURING ELEMENT - ROTATING PADDLE> is a <LEVEL MEASURING ELEMENT - MECHANICAL> that is an established technique for bulk solid point level indication. It uses a low-speed gear motor that rotates a paddle wheel. When the paddle is stalled by solid materials, the motor is rotated on its shaft by its own torque until a flange mounted on the motor contacts a mechanical switch | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - MECHANICAL |  |  |  |  | LEVEL MEASURING ELEMENT - SERVO | [5 empty] | rdl:RDS222571 |  | A <LEVEL MEASURING ELEMENT - SERVO> is a <LEVEL MEASURING ELEMENT - MECHANICAL> that uses a float and a wire fed by a servo motor to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - MECHANICAL |  |  |  |  | LEVEL MEASURING ELEMENT - VIBRATING | [5 empty] | rdl:RDS2224563 |  | A <LEVEL MEASURING ELEMENT - VIBRATING> is a <LEVEL MEASURING ELEMENT - MECHANICAL> that is vibrating to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - VIBRATING | [5 empty] | LEVEL MEASURING ELEMENT - SINGLE PROBE VIBRATING | [5 empty] | rdl:RDS222572 |  | A <LEVEL MEASURING ELEMENT - SINGLE PROBE VIBRATING> is a <LEVEL MEASURING ELEMENT - VIBRATING> for bulk powder level measurement. It is not affected by dust, static-charge build-up from dielectric powders, or changes in conductivity, temperature, pressure, humidity or moisture content. Since only one sensing element contacts the powder, bridging between two probe elements is eliminated and media build-up is minimized | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - VIBRATING | [5 empty] | LEVEL MEASURING ELEMENT - TUNING FORK VIBRATING | [5 empty] | rdl:RDS222573 |  | A <LEVEL MEASURING ELEMENT - TUNING FORK VIBRATING> is a <LEVEL MEASURING ELEMENT - VIBRATING> for bulk powder level measurement. It is not affected by dust, static-charge build-up from dielectric powders, or changes in conductivity, temperature, pressure, humidity or moisture content, but it is prone to material buildup between the tines | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - NUCLEAR | [5 empty] | rdl:RDS222574 |  | A <LEVEL MEASURING ELEMENT - NUCLEAR> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that uses the absorption of gamma radiation by process material to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - OPTICAL | [5 empty] | rdl:RDS222575 |  | A <LEVEL MEASURING ELEMENT - OPTICAL> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that uses the transmission, reflection or refraction of light by the process material to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - OPTICAL INFRARED | [5 empty] | rdl:RDS222576 |  | A <LEVEL MEASURING ELEMENT - OPTICAL INFRARED> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that is used for point level sensing of sediments, liquids with suspended solids, and liquid-liquid interfaces. These sensors sense the decrease or change in transmission of infrared light emitted from an infrared diode (LED) | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - OPTICAL LASER | [5 empty] | rdl:RDS222577 |  | A <LEVEL MEASURING ELEMENT - OPTICAL LASER> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that involves the use of a laser. Laser light reflects off most solid, liquid surfaces. The time of flight can be measured with precise timing circuitry, to determine the range or distance of the surface from the sensor | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - OPTICAL LASER |  |  |  |  | LEVEL MEASURING ELEMENT - OPTICAL REFRACTION | [5 empty] | rdl:RDS222578 |  | A <LEVEL MEASURING ELEMENT - OPTICAL REFRACTION> is a <LEVEL MEASURING ELEMENT - OPTICAL LASER> that uses the difference between the refractive index of the process fluid and the material of the sensing element to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - WAVE | [5 empty] | rdl:RDS222579 |  | A <LEVEL MEASURING ELEMENT - WAVE> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that measures the level by detecting either a propagation time or a damping of a propagation phenomenon of waves | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - FREE-SPACE RADAR | [5 empty] | rdl:RDS222580 |  | A <LEVEL MEASURING ELEMENT - FREE-SPACE RADAR> is a <LEVEL MEASURING ELEMENT - WAVE> that uses the travelling time of a radar beam to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FREE-SPACE RADAR | [5 empty] | LEVEL MEASURING ELEMENT - CONTINUOUS-WAVE RADAR | [5 empty] | rdl:RDS222581 |  | A <LEVEL MEASURING ELEMENT - CONTINUOUS-WAVE RADAR> is a <LEVEL MEASURING ELEMENT - FREE-SPACE RADAR> that uses a continuous radar wave to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FREE-SPACE RADAR | [5 empty] | LEVEL MEASURING ELEMENT - PULSED-WAVE RADAR | [5 empty] | rdl:RDS222582 |  | A <LEVEL MEASURING ELEMENT - PULSED-WAVE RADAR> is a <LEVEL MEASURING ELEMENT - FREE-SPACE RADAR> that uses a pulsed radar wave to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - GAMMA RAY | [5 empty] | rdl:RDS222583 |  | A <LEVEL MEASURING ELEMENT - GAMMA RAY> is a <LEVEL MEASURING ELEMENT - WAVE> that measures level by the attenuation of gamma rays passing through a process vessel. The technique is used to regulate the level of molten steel in a continuous casting process of steelmaking | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - GUIDED-WAVE RADAR | [5 empty] | rdl:RDS222584 |  | A <LEVEL MEASURING ELEMENT - GUIDED-WAVE RADAR> is a <LEVEL MEASURING ELEMENT - WAVE> that emits a wave that is guided by a rod, cable or rope to the surface of process material and uses the time of flight of the reflected wave to measure level | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - PULSED MICROWAVE | [5 empty] | rdl:RDS222585 |  | A <LEVEL MEASURING ELEMENT - PULSED MICROWAVE> is a <LEVEL MEASURING ELEMENT - WAVE> that using a Pulsed Reflectometry signal processing technique | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - TDR MICROWAVE | [5 empty] | rdl:RDS222586 |  | A <LEVEL MEASURING ELEMENT - TDR MICROWAVE> is a <LEVEL MEASURING ELEMENT - WAVE> using a Time-Domain Reflectometry signal processing technique which is a measurement of time of flight divided by the speed of light | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - WAVE |  |  |  |  | LEVEL MEASURING ELEMENT - ULTRASONIC | [5 empty] | rdl:RDS11468384 |  | A <LEVEL MEASURING ELEMENT - ULTRASONIC> is a <LEVEL MEASURING ELEMENT - WAVE> that is used for non-contact level sensing of liquids and bulk solids. It emits high frequency (20 kHz to 200 kHz) acoustic waves that are reflected back to and detected by the emitting transducer. Ultrasonic sensors can be designed to provide point level control, continuous monitoring or both | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | LEVEL MEASURING ELEMENT |  |  | LEVEL MEASURING ELEMENT - FLOAT | [5 empty] | rdl:RDS851039 |  | A <LEVEL MEASURING ELEMENT - FLOAT> is an <INSTRUMENTATION ITEM> and a <LEVEL MEASURING ELEMENT> that involves the actuation of a mechanical switch actuated by a float | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FLOAT |  |  |  |  | LEVEL MEASURING ELEMENT - FLOAT - MAGNET | [5 empty] | rdl:RDS222561 |  | A <LEVEL MEASURING ELEMENT - FLOAT - MAGNET> is a <LEVEL MEASURING ELEMENT - FLOAT> that uses a magnet and resistors to determine the position of the float | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FLOAT |  |  |  |  | LEVEL MEASURING ELEMENT - FLOAT - RESISTIVE CHAIN | [5 empty] | rdl:RDS222562 |  | A <LEVEL MEASURING ELEMENT - FLOAT - RESISTIVE CHAIN> is a <LEVEL MEASURING ELEMENT - FLOAT> where a permanent magnet, sealed inside a float, moves up and down a stem in which closely spaced switches and resistors are sealed. When the switches are closed, the resistance is summed and converted to current or voltage signals that are proportional to the level of the liquid | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FLOAT |  |  |  |  | LEVEL MEASURING ELEMENT - MAGNETORESISTIVE FLOAT | [5 empty] | rdl:RDS222556 |  | A <LEVEL MEASURING ELEMENT - MAGNETORESISTIVE FLOAT> is a <LEVEL MEASURING ELEMENT - FLOAT> where a permanent magnet pair is sealed inside the float arm pivot. As the float moves up the motion and location are transmitted as the angular position of the magnetic field | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FLOAT |  |  |  |  | LEVEL MEASURING ELEMENT - MAGNETOSTRICTIVE | [5 empty] | rdl:RDS222560 |  | A <LEVEL MEASURING ELEMENT - MAGNETOSTRICTIVE> is a <LEVEL MEASURING ELEMENT -> <FLOAT> that uses the magnetostrictive effect to determine the position of the float | [5 empty] | dm:ClassOfInanimatePhysicalObject
LEVEL MEASURING ELEMENT - FLOAT |  |  |  |  | LEVEL MEASURING ELEMENT - MAGNETOSTRICTIVE FLOAT | [5 empty] | rdl:RDS222557 |  | A <LEVEL MEASURING ELEMENT - MAGNETOSTRICTIVE FLOAT> is a <LEVEL MEASURING ELEMENT - FLOAT> where a permanent magnet, sealed inside a float, travels up and down a stem in which a magnetostrictive wire is sealed. For high-accuracy, continuous level measurement of a wide variety of liquids | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | MIST DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS1008899 |  | A <MIST DETECTING MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of a mist | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | MIST DETECTING MEASURING ELEMENT |  |  | OIL MIST DETECTING SENSOR | [5 empty] | rdl:RDS2225335 |  | An <OIL MIST DETECTING SENSOR> is an <INSTRUMENTATION ITEM> and an <MIST DETECTING MEASURING ELEMENT> for measurement of an <OIL MIST>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | MIST DETECTING MEASURING ELEMENT |  |  | FOG DETECTING SENSOR | [5 empty] | rdl:RDS2225336 |  | A <FOG DETECTING SENSOR> is an <INSTRUMENTATION ITEM> and an <MIST DETECTING MEASURING ELEMENT> for measurement of a atmospeheric fog. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | POSITION MEASURING ELEMENT | [5 empty] | rdl:RDS222587 |  | A <POSITION MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of physical <POSITION> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | POSITION MEASURING ELEMENT |  |  | LIMIT SWITCH SENSOR | [5 empty] | rdl:RDS1325720321 |  | A <LIMIT SWITCH SENSOR> is an <INSTRUMENTATION ITEM> and a <POSITION MEASURING ELEMENT> that is used in conjunction with a <LIMIT SWITCH> | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | PRESSURE MEASURING ELEMENT | [5 empty] | rdl:RDS462104 |  | A <PRESSURE MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <PRESSURE> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | PRESSURE MEASURING ELEMENT |  |  | ABSOLUTE PRESSURE MEASURING ELEMENT | [5 empty] | rdl:RDS222588 |  | An <ABSOLUTE PRESSURE MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PRESSURE MEASURING ELEMENT> for the measurement of <ABSOLUTE PRESSURE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
ABSOLUTE PRESSURE MEASURING ELEMENT |  |  |  |  | ABSOLUTE PRESSURE MEASURING ELEMENT - DIAPHRAGM | [5 empty] | rdl:RDS222589 |  | An <ABSOLUTE PRESSURE MEASURING ELEMENT - DIAPHRAGM> is an <ABSOLUTE PRESSURE MEASURING ELEMENT> that is a diaphragm with a reference chamber on the side of that diaphragm that not subject to pressure. In that reference chamber an absolute vacuum is given as reference pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PRESSURE MEASURING ELEMENT |  |  | DIFFERENTIAL PRESSURE MEASURING ELEMENT | [5 empty] | rdl:RDS5770524 |  | A <DIFFERENTIAL PRESSURE MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PRESSURE MEASURING ELEMENT> for the measurement of <DIFFERENTIAL PRESSURE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
DIFFERENTIAL PRESSURE MEASURING ELEMENT |  |  |  |  | DIFFERENTIAL PRESSURE MEASURING ELEMENT - BELLOWS | [5 empty] | rdl:RDS222591 |  | A <DIFFERENTIAL PRESSURE MEASURING ELEMENT - BELLOWS> is a <DIFFERENTIAL PRESSURE MEASURING ELEMENT> that is a bellows. The pressure of the media acts on both sides of this bellows upon which it moves, and this movement is a measurement for the pressure difference | [5 empty] | dm:ClassOfInanimatePhysicalObject
DIFFERENTIAL PRESSURE MEASURING ELEMENT |  |  |  |  | DIFFERENTIAL PRESSURE MEASURING ELEMENT - DIAPHRAGM | [5 empty] | rdl:RDS222590 |  | A <DIFFERENTIAL PRESSURE MEASURING ELEMENT - DIAPHRAGM> is a <DIFFERENTIAL PRESSURE MEASURING ELEMENT> where two sealed media chambers are separated by a diaphragm. The deflection caused in this way is used as a measurement for the pressure difference | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PRESSURE MEASURING ELEMENT |  |  | GAUGE PRESSURE MEASURING ELEMENT | [5 empty] | rdl:RDS1322274131 |  | A <GAUGE PRESSURE MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <PRESSURE MEASURING ELEMENT> for the measurement of <GAUGE PRESSURE> | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT |  |  |  |  | DIAPHRAGM SEAL ASSEMBLY | [5 empty] | rdl:RDS1469940001 |  | A <DIAPHRAGM SEAL ASSEMBLY> is a <GAUGE PRESSURE MEASURING ELEMENT> that consists of a <DIAPHRAGM SEAL> and additional parts like <CAPILLARY TUBE> required to transfer the pressure from the fluid to the process connection of the device in focus | [5 empty] | dm:ClassOfInanimatePhysicalObject
DIAPHRAGM SEAL ASSEMBLY | [5 empty] | TWO REMOTE SEALS DIAPHRAGM SEAL ASSEMBLY | [5 empty] | rdl:RDS1469942981 |  | A <TWO REMOTE SEALS DIAPHRAGM SEAL ASSEMBLY> is a <DIAPHRAGM SEAL ASSEMBLY> that consists of two <DIAPHRAGM SEAL>s and <CAPILLARY TUBE>s, and additional parts required for the connection | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT |  |  |  |  | GAUGE PRESSURE MEASURING ELEMENT - DIAPHRAGM | [5 empty] | rdl:RDS708344 |  | A <GAUGE PRESSURE MEASURING ELEMENT - DIAPHRAGM> is a <GAUGE PRESSURE MEASURING ELEMENT> that is a circular-shaped, corrugated membrane, either clamped around the rim between two flanges or welded and subject to the pressure of the media acting on one side. The deflection caused in this way is used as a measurement for pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT - DIAPHRAGM | [5 empty] | GAUGE PRESSURE TRANSDUCER | [5 empty] | rdl:RDS222594 |  | A <GAUGE PRESSURE TRANSDUCER> is a <GAUGE PRESSURE MEASURING ELEMENT - DIAPHRAGM> where pressure applied to a pressure transducer that produces a deflection of a <DIAPHRAGM> which introduces strain to strain gages. The strain gages produce an electrical resistance change proportional to the pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT |  |  |  |  | GAUGE PRESSURE MEASURING ELEMENT - BELLOWS | [5 empty] | rdl:RDS222592 |  | A <GAUGE PRESSURE MEASURING ELEMENT - BELLOWS> is a <GAUGE PRESSURE MEASURING ELEMENT> that is a bellows. The pressure of the media acts on the inside of this bellows upon which it moves, and this movement is a measurement for the pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT |  |  |  |  | GAUGE PRESSURE MEASURING ELEMENT - BOURDON | [5 empty] | rdl:RDS462239 |  | A <GAUGE PRESSURE MEASURING ELEMENT - BOURDON> is a <GAUGE PRESSURE MEASURING ELEMENT> that is a radially-formed tube with an oval crosssection. The pressure of the media acts on the inside of this tube. The end of the tube which is not fixed moves, and this movement is a measurement for the pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
GAUGE PRESSURE MEASURING ELEMENT |  |  |  |  | GAUGE PRESSURE MEASURING ELEMENT - CAPSULE | [5 empty] | rdl:RDS222593 |  | A <GAUGE PRESSURE MEASURING ELEMENT - CAPSULE> is a <GAUGE PRESSURE MEASURING ELEMENT> that is an element that comprises two circular-shaped, corrugated diaphragms, joined together around the rim with a pressure-tight seal. The pressure acts on the inside of this capsule and the generated stroke movement is indicated by means of a movement as a measurement of pressure | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | PROXIMITY MEASURING ELEMENT | [5 empty] | rdl:RDS850904 |  | A <PROXIMITY MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <PROXIMITY> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | PROXIMITY MEASURING ELEMENT |  |  | ACOUSTIC PROXIMITY PROBE | [5 empty] | rdl:RDS1466004581 |  | An <ACOUSTIC PROXIMITY PROBE> is an <INSTRUMENTATION ITEM> and a <PROXIMITY MEASURING ELEMENT> where acoustic signals are sent out, with the time of echo return being a measure of distance to an obstacle | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PROXIMITY MEASURING ELEMENT |  |  | CAPACITIVE PROXIMITY PROBE | [5 empty] | rdl:RDS11479520 |  | A <CAPACITIVE PROXIMITY PROBE> is an <INSTRUMENTATION ITEM> and a <PROXIMITY MEASURING ELEMENT> in which when power is applied to the sensor, an electrostatic field is generated and reacts to changes in capacitance caused by the presence of a target | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PROXIMITY MEASURING ELEMENT |  |  | INDUCTIVE PROXIMITY PROBE | [5 empty] | rdl:RDS11468260 |  | An <INDUCTIVE PROXIMITY PROBE> is an <INSTRUMENTATION ITEM> and a <PROXIMITY MEASURING ELEMENT> that senses proximity based on the use of induced magnetic fields | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | PROXIMITY MEASURING ELEMENT |  |  | INFRARED PROXIMITY PROBE | [5 empty] | rdl:RDS1466002331 |  | An <INFRARED PROXIMITY PROBE> is an <INSTRUMENTATION ITEM> and a <PROXIMITY MEASURING ELEMENT> that operates by sending out a beam of IR light, and then computing the distance to any nearby objects from characteristics of the returned (reflected) signal | [5 empty] | dm:ClassOfInanimatePhysicalObject
DETECTOR | PROXIMITY MEASURING ELEMENT |  |  | PIG DETECTOR | [5 empty] | rdl:RDS416249 |  | A <PIG INDICATOR> is a <DETECTOR> using a <PROXIMITY MEASURING ELEMENT> that is showing when/if the <PIG> is inside a <PIG RECEIVER> | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | RADIATION MEASURING ELEMENT | [5 empty] | rdl:RDS1336333241 |  | A <RADIATION MEASURING ELEMENT> is a <MEASURING ELEMENT> for detecting the level of nuclear <RADIATION> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | RADIATION MEASURING ELEMENT |  |  | GAMMA DETECTING MEASURING ELEMENT | [5 empty] | rdl:RDS880199 |  | A <GAMMA DETECTING MEASURING ELEMENT> is an <INSTRUMENTATION ITEM> and a <RADIATION MEASURING ELEMENT> for measurement of the level of gamma radiation | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | ROTATIONAL SPEED MEASURING ELEMENT | [5 empty] | rdl:RDS5773110 |  | A <ROTATIONAL SPEED MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <ROTATIONAL SPEED> | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | SAND DETECTION MEASURING ELEMENT | [5 empty] | rdl:RDS5781223 |  | A <SAND DETECTION MEASURING ELEMENT> is a <MEASURING ELEMENT> for sand detection in a pipeline | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | SMOKE DETECTION MEASURING ELEMENT | [5 empty] | rdl:RDS5781225 |  | A <SMOKE DETECTION MEASURING ELEMENT> is a <MEASURING ELEMENT> for the detection of smoke | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT | [5 empty] | rdl:RDS445454 |  | A <TEMPERATURE MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <TEMPERATURE> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT - BIMETAL | [5 empty] | rdl:RDS445184 |  | A <TEMPERATURE MEASURING ELEMENT - BIMETAL> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that is composed of two separate metals joined together and which convert a temperature change into mechanical displacement | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT - FILLED BULB | [5 empty] | rdl:RDS1029959 |  | A <TEMPERATURE MEASURING ELEMENT - FILLED BULB> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that uses the principle of fluid expansion to measure temperature | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT - INFRARED | [5 empty] | rdl:RDS222595 |  | A <TEMPERATURE MEASURING ELEMENT - INFRARED> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that infers temperature from a portion of the thermal radiation sometimes called blackbody radiation emitted by the object being measured | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT - RESISTANCE | [5 empty] | rdl:RDS445274 |  | A <TEMPERATURE MEASURING ELEMENT - RESISTANCE> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that is a sensing element that operate on the principle of changes in electrical resistance of pure metals (sintered semiconductor material) and are characterized by an almost linear positive change in resistance with temperature | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - RESISTANCE |  |  |  |  | NEGATIVE RESISTANCE TEMPERATURE ELEMENT | [5 empty] | rdl:RDS445364 |  | A <NEGATIVE RESISTANCE TEMPERATURE ELEMENT> is a <TEMPERATURE MEASURING ELEMENT - RESISTANCE> that has a negative temperature coefficient which means that the electrical resistance decreases with increasing temperature | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - RESISTANCE |  |  |  |  | POSITIVE RESISTANCE TEMPERATURE ELEMENT | [5 empty] | rdl:RDS445319 |  | A <POSITIVE RESISTANCE TEMPERATURE ELEMENT> is a <TEMPERATURE MEASURING ELEMENT - RESISTANCE> that has a positive temperature coefficient which means that the electrical resistance increases with increasing temperature | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - RESISTANCE |  |  |  |  | PT100 RESISTANCE THERMOMETER | [5 empty] | rdl:RDS890324 |  | A <PT100 RESISTANCE THERMOMETER> is a <TEMPERATURE MEASURING ELEMENT - RESISTANCE> made of platinum that has the resistance of 100 ohm at 0 deg. Celsius | [5 empty] | dm:ClassOfInanimatePhysicalObject
PT100 RESISTANCE THERMOMETER | [5 empty] | WINDING TEMPERATURE SENSOR | [5 empty] | rdl:RDS1186559 |  | A <WINDING TEMPERATURE SENSOR> is a <PT100 RESISTANCE THERMOMETER> that is designed for winding temperature measurement | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - RESISTANCE |  |  |  |  | PT1000 RESISTANCE THERMOMETER | [5 empty] | rdl:RDS1464287521 |  | A <PT1000 RESISTANCE THERMOMETER> is a <TEMPERATURE MEASURING ELEMENT - RESISTANCE> made of platinum that has the resistance of 1000 ohm at 0 deg. Celsius | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - RESISTANCE |  |  |  |  | THERMISTOR | [5 empty] | rdl:RDS6624243 |  | A <THERMISTOR> is a <TEMPERATURE MEASURING ELEMENT - RESISTANCE> that use a semiconductor device to change resistance that increases with increasing temperature, and is used for the sensitive measurement and control | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE | [5 empty] | rdl:RDS445229 |  | A <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that is a temperature sensing element in which a pair of wires of dissimilar metals (as copper and iron) are joined and the free ends of the wires are connected to an instrument (as a voltmeter) that measures the difference in potential created at the junction of the two metals | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE B THERMOCOUPLE | [5 empty] | rdl:RDS12171850 |  | A <TYPE B THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 70%Pt/30%Rh and (-) 94%Pt/6%Rh, by weight | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE C THERMOCOUPLE | [5 empty] | rdl:RDS12171895 |  | A <TYPE C THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 95%W/5%Re and (-) 74%W/26%Re, by weight | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE D THERMOCOUPLE | [5 empty] | rdl:RDS12171940 |  | A <TYPE D THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 97%W/3%Re and (-) 75%W/25%Re, by weight | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE E THERMOCOUPLE | [5 empty] | rdl:RDS12171985 |  | A <TYPE E THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) Ni-Cr (chromel) and (-) Cu-Ni (constantan). | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE G THERMOCOUPLE | [5 empty] | rdl:RDS12172030 |  | A <TYPE G THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) W and (-) 74%W/26%Re, by weight | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE J THERMOCOUPLE | [5 empty] | rdl:RDS12172075 |  | A <TYPE J THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made (+) Fe and (-) Cu-Ni (constantan). | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE K THERMOCOUPLE | [5 empty] | rdl:RDS12171355 |  | A <TYPE K THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made (+) Ni-Cr (chromel) and (-) Ni-Al (alumel). | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE M THERMOCOUPLE | [5 empty] | rdl:RDS222596 |  | A <TYPE M THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 82%Ni/18%Mo and (-) 99.2%Ni/0.8%Co%, by weight. | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE N THERMOCOUPLE | [5 empty] | rdl:RDS222597 |  | A <TYPE N THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 14.4%Cr/1.4%Si/0.1%Mg (Nicrosil) and (-) 95.5%Ni/4.4%Si/0.1%Mg (Nisil). | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE R THERMOCOUPLE | [5 empty] | rdl:RDS12170905 |  | A <TYPE R THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 87%Pt/13%Rh by weight and (-) Pt. | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE S THERMOCOUPLE | [5 empty] | rdl:RDS12170950 |  | A <TYPE S THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) 90%Pt/10%Rh by weight and (-) Pt. | [5 empty] | dm:ClassOfInanimatePhysicalObject
TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE |  |  |  |  | TYPE T THERMOCOUPLE | [5 empty] | rdl:RDS12170995 |  | A <TYPE T THERMOCOUPLE> is a <TEMPERATURE MEASURING ELEMENT - THERMOCOUPLE> made of (+) Cu and (-) Cu-Ni (constantan). | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | TEMPERATURE MEASURING ELEMENT |  |  | THERMAL ELEMENT BULB | [5 empty] | rdl:RDS7350739 |  | A <THERMAL ELEMENT BULB> is an <INSTRUMENTATION ITEM> and a <TEMPERATURE MEASURING ELEMENT> that is a closed metal bulb containing a temperature responsive fluid | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | VELOCITY MEASURING ELEMENT | [5 empty] | rdl:RDS54808340199 |  | A <VELOCITY MEASURING ELEMENT> is a <MEASURING ELEMENT> that is sensing velocity (vibration) conforming ISO2041 | [5 empty] | dm:ClassOfFunctionalObject
VELOCITY MEASURING ELEMENT | INSTRUMENTATION ITEM |  |  | VELOCITY PICKUP | [5 empty] | rdl:RDS5480833728 |  | A <VELOCITY PICKUP> is an <INSTRUMENTATION ITEM> and a <VELOCITY MEASURING ELEMENT> that uses the principle set forth in ISO 2041 | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | VIBRATION MEASURING ELEMENT | [5 empty] | rdl:RDS435959 |  | A <VIBRATION MEASURING ELEMENT> is a <MEASURING ELEMENT> that senses vibration | [5 empty] | dm:ClassOfFunctionalObject
MEASURING ELEMENT |  |  | WATER HARDNESS MEASURING ELEMENT | [5 empty] | rdl:RDS2229120 |  | A <WATER HARDNESS MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <HARDNESS> of <WATER> | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | WATER HARDNESS MEASURING ELEMENT |  |  | WATER HARDNESS SENSOR | [5 empty] | rdl:RDS2225337 |  | A <WATER HARDNESS SENSOR> is an <INSTRUMENTATION ITEM> and a <WATER HARDNESS MEASURING ELEMENT>. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | HEIGHT MEASURING ELEMENT | [5 empty] | rdl:RDS2225338 |  | A <HEIGHT SENSOR> is a <MEASURING ELEMENT> that measures <HEIGHT>. | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | HEIGHT MEASURING ELEMENT |  | WAVE HEIGHT SENSOR | [5 empty] | rdl:RDS1300144801 |  | A <WAVE HEIGHT SENSOR> is an <INSTRUMENTATION ITEM> and a <HEIGHT MEASURING ELEMENT> that senses wave heights and produces an output signal to represent the measured height. | [5 empty] | dm:ClassOfInanimatePhysicalObject
MEASURING ELEMENT |  |  | WEIGHT MEASURING ELEMENT | [5 empty] | rdl:RDS222599 |  | A <WEIGHT MEASURING ELEMENT> is a <MEASURING ELEMENT> for the measurement of <WEIGHT> (mass) | [5 empty] | dm:ClassOfFunctionalObject
INSTRUMENTATION ITEM | WEIGHT MEASURING ELEMENT |  |  | CAPACITIVE LOAD CELL | [5 empty] | rdl:RDS222600 |  | A <CAPACITIVE LOAD CELL> is an <INSTRUMENTATION ITEM> and a <WEIGHT MEASURING ELEMENT where the capacitance of a capacitor changes as the load presses the two plates of a capacitor closer together | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | WEIGHT MEASURING ELEMENT |  |  | HYDRAULIC LOAD CELL | [5 empty] | rdl:RDS222601 |  | A <HYDRAULIC LOAD CELL> is an <INSTRUMENTATION ITEM> and a <WEIGHT MEASURING ELEMENT that is a conventional piston and cylinder arrangement | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | WEIGHT MEASURING ELEMENT |  |  | PIEZOELECTRIC LOAD CELL | [5 empty] | rdl:RDS222602 |  | A <PIEZOELECTRIC LOAD CELL> is an <INSTRUMENTATION ITEM> and a <WEIGHT MEASURING ELEMENT that works on the same principle of deformation as the strain gauge load cells, but a voltage output is generated by the basic piezoelectric material - proportional to the deformation of load cell | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | WEIGHT MEASURING ELEMENT |  |  | PNEUMATIC LOAD CELL | [5 empty] | rdl:RDS222603 |  | A <PNEUMATIC LOAD CELL> is an <INSTRUMENTATION ITEM> and a <WEIGHT MEASURING ELEMENT that has a diaphragm working on a force-balance principle | [5 empty] | dm:ClassOfInanimatePhysicalObject
INSTRUMENTATION ITEM | WEIGHT MEASURING ELEMENT |  |  | VIBRATING WIRE LOAD CELL | [5 empty] | rdl:RDS222604 |  | A <VIBRATING WIRE LOAD CELL> is an <INSTRUMENTATION ITEM> and a <WEIGHT MEASURING ELEMENT that measures force (a mechanical load) using a wire that vibrates at a high frequency. The applied external force changes the tension on the wire, this changes the frequency | [5 empty] | dm:ClassOfInanimatePhysicalObject